## Slide 1: Microwave Remote Sensing of the Boundary Layer
Harald Czekala1, Oscar Hartogensis2, Martin Philipp1, Thomas Rose1
1: RPG Radiometer Physics GmbH, Meckenheim, Germany
2: Wageningen University, Wageningen, The Netherlands

## Slide 2: Outline
Passive microwave radiometer for Boundary Layer (BL) T-profiling:
  The concept of elevation scans with saturated channels
  The expected signal
  Requirements on instrument design / receiver improvement
Active / Scintillometry:
  Concept of direct measurement of latent heat flux (evapotranspiration) (when combined with optical scintillometer)
  Design of Transmit/Receive system at 160.8 GHz
  Prototype assembly
Active / FMCW cloud radar at 94 GHz
  Concept, design drivers
  Preliminary Instrument specifications
Summary / Conclusions

## Slide 3: µ-wave BL-Temperature Profiling: The Concept
Absorption length decreases for frequencies at line center (60 GHz)
Limited (approx. 600 m) emission depth @ 58 GHz
Weighting functions shifted towards surface by elevation scans

## Slide 4: Simulated TB Signature of BL-T-profiles
Small TB variation
on top of large signal
Low-Level inv.:
clear signal
curvature)
Lifted inv. difficult
(parallel lines)
 Channels with lower saturation
2.5 K signal
1.0 K signal

## Slide 5: µ-wave BL-Temperature Profiling: The Instrument
HATPRO Direct Detection Filterbank Profiler (parallel data aquisition 14 ch.)
7-channel
V-Band
Receiver

## Slide 6: MWR Design Driver: Noise Reductionat 58 GHz
Courtesy of Gerrit Maschwitz, University of Cologne, Germany
To resolve shape of temperature profile, the radiometer needs
Increased bandwidth at saturated channels only (2000 MHz vs. 230 MHz, noise reduction 1:3)
Increased integration time: 15 s
Multi-channel obs. (in parallel) to avoid saturation (limited range)
Large optics, small beam
Recent developments:
  Space qualified/optimised RF amplifiers from UMS
  Better noise temperature
  Better 1/f performance (MMIC material GaAs vs. InP)
  Better calibration repeatability (standing wave problem solved)
RMS TB58 < 0.05 K
1.8°
300 K full system

## Slide 7: Result: Time Series of Temperature Inversion
Regr. | 90° | 30° | 19.2° | 14.4° | 11.4° | 8.4° | 6.6° | 4.8°
Linear | 1 – 7 | 4 – 7 | 4 – 7 | 6 + 7 | 6 + 7 | 7 | 7 | 7
Quad. | 1 – 7 | – | – | – | – | – | – | –
Airm. | 1.00 | 2.00 | 3.04 | 4.02 | 5.05 | 6.84 | 8.70 | 11.95
Regression Retrievals: Elevation angles, airmass, and channel usage

## Slide 8: Comparison with Mast data: 15 day time series
Courtesy of Susanne Crewell, University of Cologne, Germany
10 m temperature:
Black: mast, red: MWR
100 m temperature:
Black: mast, red: MWR
Temp. Gradient: show
lapse rate vs. inversions

## Slide 9
RPG-MMWS-160
A Millimeter-wave Scintillometer
for Latent Heat Flux Measurements
  Co-operation of Wageningen University (WU) and RPG
  Wageningen: Concept, user requirements, test, evaluation
  RPG: Technical design and realisation
  Prototype: Jointly funded by STW, The Netherlands, and RPG
  First instrument to be tested in September/October 2012

## Slide 10: Scintillometry : The Concept

## Slide 11: Scintillometry: The Concept / Theory
λ2=1.86 mm
λ1=930 nm
Monin-Obukhov
Similarity Theory

## Slide 12: Scintillometer Design
Key features of the MWSC:
Frequency: 160.8 GHz (λ=1.86 mm) for small beam and sufficient co-spectrum with optical LAS
300 mm Cassegrain optics (51 dB gain)
40 mW transmitted power
Detection bandwidth: 10 kHz
Sensitivity of 90 dB

## Slide 13: Scintillometer Technology
MWSC Technology Heritage:
Housing, power supply, T-control: Heritage from µ-wave radiometer
160 GHz source originally designed as 166 GHz source for MetOp-SG 664 LO
Receiver: Modified 150 GHz Space FE
Source stability (with AGC): 1.8 e-6
Ultra-stable frequency at Tx and Rx units
DC
AC

## Slide 14: MWSC Integration: End of August 2012

## Slide 15: MWSC Integration: End of August 2012

## Slide 16
RPG-FMCW-94
RPG-FMCW-94-DP
A Cloud Radar by RPG
  Frequency Modulated, Continuous Wave, 94 GHz
  Low mode for fog, high mode up to 12 km
  Re-using RPG technology from Tx/Rx systems and network analysis
  Re-using instrument infrastructure (housing, steering, control, data processing) from microwave radiometers
  Supported by Fraunhofer Institute for RF Technology and Radar (FhG-FHR, Wachtberg, Germany, formerly FGAN)
  Availability: Mid of 2013

## Slide 17: FMCW Radar Concept (in brief)
Distance to Object:
Altitude Resolution:
Radar Reflectivity:
Velocity:

## Slide 18: RPG-FMCW-94 Specifications (I)
Operating Frequency: 94 GHz +/-150 MHz
IF Range: 0.5 MHz to 1.3 MHz
Continuous Power: 500 mW (Solid State)
T/R Type: Bi-static
Antenna Diameter: 500 mm
Gain: 52 dB
Chirp Rate: 100 / sec
Chirp Variations: 7
Passive LWP Channel: 89 Ghz (optional)
Dynamic Range: -100 dBz to +20 dBz

## Slide 19: RPG-FMCW-94 Specifications (II)
Ranging: 10 m to 12 km
Sampling Rate: 1 / 4 seconds
Vertical Resolution: 1 m (r: 10m - 600 m)
2 m (r: 0.6 – 1.0 km)
4 m (r: 1.0 – 2.5 km)
8 m (r: 2.5 – 5.0 km)
16 m (r: 5.0 – 12.0 km)
Doppler Resolution: +/- 15 cm/sec
Polarisation: v / h (optional)
  Availability: Mid of 2013
  Possible extensions by
  Additional frequencies
  Passive microwave channels
  Polarisation options

## Slide 20: Summary / Conclusions
Microwave Radiometers
  High-resolution boundary layer tenperature profiling possible with suitable/optimized receiver architecture
  Optimization / synergy with space projects (60 GHz LNA, calibration concepts)
Scintillation
  Transfer of known concept (optical LAS for sensoble heat flux) to microwave / millimeter wave region @160 GHz
FMCW cloud radar
  Transfer of existing Tx/Rx technology from Lab to field
Rapid Prototyping and reduced time to market only possible due to broad field of applications at RPG.

## Slide 21
Thank you!
czekala@radiometer-physics.de